## Data
start_row | 5
stop_row | 40
V1_set | V2_set | n | Start/xl del. | AZ1 del. | Range del. | V2 |  |  | Vd1 |  |  | V1 |  |  | dvm_T1 | dvm_T2 | GMH_T1 | GMH_T2 | Ambient conditions |  |  | Comment | Role | Instrument descr.
Run Id: | HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 |  |  |  |  | t | V | sd(V) | t | V | sd(V) | t | V | sd(V) |  |  | 20.79 | 20.92 | T | P(mbar) | RH(%)
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 11:42:16 | -1.0003 | 7.58e-07 | 03/03/2017 11:44:14 | -0.00030452 | 7.4346e-07 | 03/03/2017 11:40:36 | 10 | 4.9264e-07 | 108.0052 | 9978.7676 | 20.79 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | SRC2 | SRC: F5520A
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 11:49:10 | 1.0003 | 7.3424e-07 | 03/03/2017 11:51:08 | 0.0003035 | 5.5713e-07 | 03/03/2017 11:47:30 | -10 | 3.7702e-07 | 108.0007 | 9978.7739 | 20.77 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | DVMT1 | none
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 11:56:04 | -1.0003 | 5.9878e-07 | 03/03/2017 11:58:02 | -0.00030446 | 6.3571e-07 | 03/03/2017 11:54:25 | 10 | 4.7347e-07 | 107.996 | 9978.7739 | 20.77 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | SRC1 | SRC: D4808
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 12:02:57 | -1.0103 | 3.1889e-07 | 03/03/2017 12:04:55 | -0.0093983 | 4.6224e-07 | 03/03/2017 12:01:18 | 10 | 4.2956e-07 | 108.0169 | 9978.7782 | 20.77 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | GMH1 | GMH: s/n628
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:09:50 | -1.0003 | 7.4657e-07 | 03/03/2017 12:11:48 | -0.0003045 | 9.2626e-07 | 03/03/2017 12:08:11 | 10 | 3.0948e-07 | 107.9939 | 9978.7782 | 20.78 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | GMH2 | GMH: s/n627
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:16:44 | 1.0003 | 5.1851e-07 | 03/03/2017 12:18:42 | 0.00030413 | 7.3749e-07 | 03/03/2017 12:15:05 | -10 | 3.593e-07 | 107.9999 | 9978.7782 | 20.8 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | DVMT2 | DVM: HP34401A, s/n976
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:23:38 | -1.0003 | 3.5549e-07 | 03/03/2017 12:25:36 | -0.0003052 | 5.2179e-07 | 03/03/2017 12:21:59 | 10 | 5.1713e-07 | 108.0014 | 9978.7718 | 20.81 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | DVMd | DVM: HP3458A, s/n452
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 12:30:31 | -1.0103 | 5.625e-07 | 03/03/2017 12:32:29 | -0.0093989 | 7.5929e-07 | 03/03/2017 12:28:52 | 10 | 4.5726e-07 | 108.0025 | 9978.7739 | 20.82 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | switchbox | V1
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:37:24 | -1.0003 | 6.355e-07 | 03/03/2017 12:39:22 | -0.00030512 | 6.3054e-07 | 03/03/2017 12:35:45 | 10 | 4.2515e-07 | 107.9959 | 9978.7697 | 20.83 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | DVM12 | DVM: HP3458A, s/n518
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:44:18 | 1.0003 | 5.4757e-07 | 03/03/2017 12:46:16 | 0.00030431 | 5.218e-07 | 03/03/2017 12:42:39 | -10 | 5.9233e-07 | 108.0173 | 9978.7739 | 20.84 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | GMHroom | GMH: s/n367
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 12:51:12 | -1.0003 | 7.0018e-07 | 03/03/2017 12:53:10 | -0.00030492 | 7.5731e-07 | 03/03/2017 12:49:33 | 10 | 4.0343e-07 | 107.9879 | 9978.7761 | 20.83 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 12:58:05 | -1.0103 | 9.1065e-07 | 03/03/2017 13:00:03 | -0.0093989 | 6.921e-07 | 03/03/2017 12:56:26 | 10 | 3.8119e-07 | 107.9968 | 9978.7803 | 20.84 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:04:58 | -1.0003 | 7.2687e-07 | 03/03/2017 13:06:56 | -0.00030506 | 7.0195e-07 | 03/03/2017 13:03:19 | 10 | 4.8661e-07 | 107.9922 | 9978.7803 | 20.83 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:11:52 | 1.0003 | 5.4282e-07 | 03/03/2017 13:13:50 | 0.00030448 | 5.9097e-07 | 03/03/2017 13:10:13 | -10 | 2.081e-07 | 107.9867 | 9978.7803 | 20.83 | 20.92 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:18:46 | -1.0003 | 7.1355e-07 | 03/03/2017 13:20:44 | -0.00030495 | 7.7584e-07 | 03/03/2017 13:17:07 | 10 | 5.6858e-07 | 108.0136 | 9978.7761 | 20.83 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 13:25:39 | -1.0103 | 6.0474e-07 | 03/03/2017 13:27:37 | -0.0093983 | 5.535e-07 | 03/03/2017 13:24:00 | 10 | 2.4927e-07 | 108.0003 | 9978.7803 | 20.83 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:32:32 | -1.0003 | 4.7052e-07 | 03/03/2017 13:34:30 | -0.00030469 | 6.6559e-07 | 03/03/2017 13:30:53 | 10 | 5.474e-07 | 108.0064 | 9978.7761 | 20.82 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:39:26 | 1.0003 | 7.8901e-07 | 03/03/2017 13:41:24 | 0.00030443 | 5.9264e-07 | 03/03/2017 13:37:47 | -10 | 4.814e-07 | 108.0087 | 9978.7739 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 13:46:20 | -1.0003 | 6.1303e-07 | 03/03/2017 13:48:18 | -0.00030478 | 7.6309e-07 | 03/03/2017 13:44:41 | 10 | 5.8786e-07 | 108.003 | 9978.7739 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 13:53:13 | -1.0103 | 4.8705e-07 | 03/03/2017 13:55:11 | -0.0093985 | 6.7224e-07 | 03/03/2017 13:51:34 | 10 | 3.533e-07 | 107.999 | 9978.7655 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:00:06 | -1.0003 | 9.3526e-07 | 03/03/2017 14:02:04 | -0.00030485 | 7.3558e-07 | 03/03/2017 13:58:27 | 10 | 3.6152e-07 | 107.9969 | 9978.7697 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:07:00 | 1.0003 | 6.4054e-07 | 03/03/2017 14:08:58 | 0.0003045 | 5.9823e-07 | 03/03/2017 14:05:21 | -10 | 4.6964e-07 | 108.004 | 9978.7527 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:13:54 | -1.0003 | 8.1287e-07 | 03/03/2017 14:15:52 | -0.0003047 | 8.3014e-07 | 03/03/2017 14:12:15 | 10 | 4.6048e-07 | 108.0108 | 9978.7464 | 20.81 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 14:20:47 | -1.0103 | 8.2486e-07 | 03/03/2017 14:22:45 | -0.0093986 | 6.0389e-07 | 03/03/2017 14:19:08 | 10 | 4.8943e-07 | 107.9927 | 9978.74 | 20.81 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:27:40 | -1.0003 | 4.4525e-07 | 03/03/2017 14:29:38 | -0.00030464 | 5.5503e-07 | 03/03/2017 14:26:01 | 10 | 4.2742e-07 | 108.004 | 9978.7337 | 20.81 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:34:34 | 1.0003 | 8.3695e-07 | 03/03/2017 14:36:32 | 0.00030437 | 7.2153e-07 | 03/03/2017 14:32:55 | -10 | 4.4739e-07 | 107.9986 | 9978.7273 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:41:29 | -1.0003 | 8.4557e-07 | 03/03/2017 14:43:27 | -0.00030459 | 3.8377e-07 | 03/03/2017 14:39:49 | 10 | 3.384e-07 | 108.0079 | 9978.7294 | 20.8 | 20.91 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 14:48:22 | -1.0103 | 5.0489e-07 | 03/03/2017 14:50:20 | -0.0093982 | 4.038e-07 | 03/03/2017 14:46:42 | 10 | 3.2766e-07 | 107.994 | 9978.7167 | 20.8 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 14:55:15 | -1.0003 | 7.0761e-07 | 03/03/2017 14:57:13 | -0.00030464 | 9.4841e-07 | 03/03/2017 14:53:35 | 10 | 5.3141e-07 | 107.9956 | 9978.7082 | 20.79 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 15:02:09 | 1.0003 | 4.6451e-07 | 03/03/2017 15:04:07 | 0.000304 | 3.5423e-07 | 03/03/2017 15:00:29 | -10 | 5.1217e-07 | 108.0043 | 9978.6955 | 20.77 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 15:09:03 | -1.0003 | 7.2059e-07 | 03/03/2017 15:11:01 | -0.00030454 | 3.8355e-07 | 03/03/2017 15:07:24 | 10 | 6.1567e-07 | 107.9964 | 9978.6828 | 20.77 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 15:15:56 | -1.0103 | 6.2834e-07 | 03/03/2017 15:17:54 | -0.0093984 | 8.4123e-07 | 03/03/2017 15:14:17 | 10 | 2.5713e-07 | 107.9703 | 9978.6743 | 20.78 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 15:22:49 | -1.0003 | 8.9769e-07 | 03/03/2017 15:24:47 | -0.00030421 | 7.7734e-07 | 03/03/2017 15:21:10 | 10 | 4.0994e-07 | 107.9916 | 9978.6488 | 20.77 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
-10 | 1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 15:29:43 | 1.0003 | 6.951e-07 | 03/03/2017 15:31:41 | 0.00030378 | 7.6338e-07 | 03/03/2017 15:28:04 | -10 | 3.7879e-07 | 108.0148 | 9978.6425 | 20.78 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0004 | 10 | 60 | 5 | 90 | 03/03/2017 15:36:37 | -1.0003 | 9.1961e-07 | 03/03/2017 15:38:35 | -0.0003045 | 5.1132e-07 | 03/03/2017 15:34:58 | 10 | 4.249e-07 | 107.995 | 9978.6382 | 20.78 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
10 | -1.0104 | 10 | 60 | 5 | 90 | 03/03/2017 15:43:30 | -1.0103 | 8.2194e-07 | 03/03/2017 15:45:28 | -0.0093981 | 4.7101e-07 | 03/03/2017 15:41:51 | 10 | 3.833e-07 | 108.0041 | 9978.6255 | 20.77 | 20.9 |  |  |  | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
Note on A1: Tim Lawson: Make sure these are updated before running HRBC or HRBA! Otherwise the wrong settings will be run or analysed.
Note on A2: Tim Lawson: Make sure these are updated before running HRBC or HRBA! Otherwise the wrong settings will be run or analysed.
Note on A4: Tim Lawson: The Run Id identifies a particular physical setup and its only use is to allow the analysis program (HRBA) to match measurement data with corresponding Rlink data. Each Run Id must be unique and is automatically generated by combining run information such as: the version of HRBC used to acquire the data; the names of the two resistors used and the date and time when the run was started. The Run Id should NOT be confused with the comment, which is intended to record and describe (as opposed to identify) the physical setup, along with any other relevant observations. The comment is intended to be read by a human, not HRBA.

## Rlink
Start row | 28 | ¦V1_set¦ | 10
N_Reversals | 10 | ¦V2_set¦ | 1
N_Readings | 10
Run Id: | HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33
Comment | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing. | [8 empty] | I_nominal
06/03/2017 10:33:16 |  | Nom. value | ¦V¦ | [6 empty] | 0.0001
R1 | Ti612 100k | 100000 | 10
R2 | SR104r 10k | 10000 | 1
ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+_av | ΔV-_av
-2.5508e-07 | -1.3577e-06 | 7.7893e-08 | -1.5265e-06 | -9.0641e-08 | -8.5265e-07 | -4.5403e-07 | -1.1512e-06 | -9.3348e-07 | -2.5551e-07 | -2.4257e-08 | -1.0477e-06
-2.3298e-07 | -8.848e-07 | 4.6963e-07 | -1.1438e-06 | 2.3672e-07 | -9.0812e-07 | -4.6877e-07 | -1.0724e-06 | 1.5667e-07 | -1.1489e-06 | 3.9183e-07 | 4.3763e-07
3.1379e-08 | -8.8662e-07 | -3.2811e-07 | -1.0835e-06 | -2.7552e-07 | -1.8879e-06 | 6.1374e-07 | -1.5178e-06 | 3.3649e-07 | -1.241e-06 | 49 | 49
-3.4041e-07 | -4.5753e-07 | 3.3335e-07 | -6.7702e-07 | -1.4015e-07 | -1.1723e-06 | -1.8197e-07 | -1.0005e-06 | 2.7491e-07 | -5.7806e-07
1.3132e-07 | -3.1457e-07 | 1.8785e-07 | -1.1138e-06 | -4.662e-07 | -1.4292e-06 | 3.2799e-07 | -1.0956e-06 | -5.4079e-07 | -5.7905e-07
2.6953e-07 | -1.2124e-06 | -7.9743e-07 | -3.648e-07 | -3.0417e-07 | -6.0672e-07 | 3.2218e-07 | -1.2665e-06 | 4.6933e-07 | -7.6162e-07 | Rd
-1.576e-08 | -2.8323e-07 | -2.1856e-07 | -1.663e-06 | 6.065e-07 | -1.0702e-06 | 1.9884e-07 | -6.167e-07 | -7.8463e-07 | -1.3153e-06 | 0.0051171
-1.2309e-07 | -1.2379e-06 | 6.4553e-07 | -2.3576e-06 | 6.7141e-08 | -8.065e-07 | -4.3257e-08 | -1.2472e-06 | 2.63e-07 | -1.6592e-06 | 0.0029371
-2.0469e-08 | -1.2239e-06 | -1.1623e-08 | -1.6839e-06 | -2.5092e-07 | -1.2076e-06 | -3.2335e-07 | -5.5927e-07 | 8.5402e-09 | -1.0974e-06 | 96.8266
5.0666e-07 | -9.176e-07 | 7.4889e-07 | -1.298e-06 | -3.686e-07 | -1.0863e-06 | 1.6639e-08 | -1.2377e-06 | -5.4359e-07 | -2.6676e-07
[12 empty] | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
[12 empty] | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
Note on A1: t.lawson: 1st row of actual data (after 6 lines of header)
Note on C1: Tim Lawson: Make sure these are updated to reflect nominal resister ratio
Note on C2: Tim Lawson: Make sure these are updated to reflect nominal resister ratio

## Parameters
Resistor Info: | [6 empty] | Instrument Info: | [6 empty] | Notes:
description | parameter | value | uncert | dof | label | Comment / Reference | description | parameter | value | uncert | dof | label | Comment / Reference | 1. Units of alpha and beta are [deg C]^-1 and [def C]^-2, respectively (i.e. 1e-6*ppm/C...).
Pt 100r | R0_LV | 100.0118 | 0 | inf | Pt_R0 | Don't use for absolute temperature measurement - differences only! | GMH: s/n367 | T_correction |  |  |  | GMH367T_cor | (Cal report?) | 2. Units of gamma are [V]^-1 (i.e. 1e-6*ppm/V).
Pt 100r | TRef_LV | 0 | 0 | inf | Pt_Tref_LV |  | GMH: s/n367 | P_correction |  |  |  | GMH367P_cor | (Cal report?) | 3. Uncertainties in Rs propagate through to dependant Rx values through the build-up. This is evident in the increasing dof as Rx increases.
Pt 100r | alpha | 0.0039091 | 0 | inf | Pt_alpha |  | GMH: s/n367 | RH_correction |  |  |  | GMH367RH_cor | (Cal report?) | 4. For the Hamon, the voltage coefficient is the effect as it [is/would be] measured on the complete configuration of elements;
Pt 100r | beta | -5.926e-07 | 0 | inf | Pt_beta |  | GMH: s/n367 | addr | 6 |  |  |  | Windows-assigned port to 3100N USB connector | The series configuration has 1/10 the voltage drop per element, compared with the parallel, for the same overall ppm change in resistance.
Pt 100r | date | 1899-12-31 |  |  |  |  | GMH: s/n367 | str_addr | COM6 |  |  |  | Windows-assigned port to 3100N USB connector | 5. Calibration dates relate to the LV measurement.
Pt 100r | T_sensor | none |  |  |  |  | GMH: s/n367 | role | UNUSED GMH, s/n367 |  |  |  |  | This shouldn't cause a problem as the LV and HV meaurements are interleaved and their means are not likely to differ by more than a few minutes.
SR104t 10k | R0_LV | 10000.2 | 0 | inf | SR104t_R0 |  | GMH: s/n367 | test | RH |  |  |  |  | 6. Measurement equation(s):
SR104t 10k | TRef_LV | 23 | 0 | inf | SR104t_Tref_LV |  | GMH: s/n627 | T_correction | -0.001 | 0.00045 | inf | GMH627_cor | (Cal report?) |  | R1 = -(R2*(1+vrc)*V1_av*G)/( G*V2_av -Vd_av)
SR104t 10k | alpha | 0.001 | 0 | inf | SR104t_alpha |  | GMH: s/n627 | addr | 5 |  |  |  | Windows-assigned port to 3100N USB connector |  | where: | R2 = R2_0*(1+R2alpha*dT2 + R2beta*dT2**2 + R2gamma*dV2) + Rd
SR104t 10k | beta | 0 | 0 | inf | SR104t_beta |  | GMH: s/n627 | str_addr | COM5 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | where: | dT2 = T2_av - R2TRef + T_def2
SR104t 10k | date | 1899-12-31 |  |  |  |  | GMH: s/n627 | role | UNUSED GMH, s/n627 | [7 empty] | ( where: | T2_av = T2_av_gmh )
SR104t 10k | T_sensor | none |  |  |  |  | GMH: s/n627 | test | T |  |  |  | Test Temperature measurement function |  |  |  | dV2 = abs(abs(V2av) - R2VRef)
SR104r 10k | R0_LV | 9999.9993 | 0.00033 | 16 | SR104r_R0_LV |  | GMH: s/n628 | T_correction | -0.0005 | 0.0004 | inf | GMH628_cor | (Cal report?) |  |  | V1av = (V1[0]-2*V1[1]+V1[2])/4
SR104r 10k | TRef_LV | 20 | 0 | inf | SR104r_Tref_LV |  | GMH: s/n628 | addr | 4 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | V2av = (V2[0]-2*V2[1]+V2[2])/4
SR104r 10k | VRef_LV | 5 | 0 | inf | SR104r_Vref_LV |  | GMH: s/n628 | str_addr | COM4 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | Vdav = (Vd[0]-2*Vd[1]+Vd[2])/4 + Vlin_Vd + Vdrift_Vd
SR104r 10k | R0_HV | 9999.9993 | 0.00033 | 16 | SR104r_R0_HV |  | GMH: s/n628 | role | UNUSED GMH, s/n628 | [6 empty] | where: | Vlin_Vd is the linearity of DVMd
SR104r 10k | TRef_HV | 20 | 0 | inf | SR104r_Tref_HV |  | GMH: s/n628 | test | T |  |  |  | Test Temperature measurement function |  |  |  | Vdrift_Vd is the zero-valued type-B drift of DVMd
SR104r 10k | VRef_HV | 10 | 0 | inf | SR104r_Vref_HV |  | V1 | addr | 2 | [6 empty] | G = (Vd[3]-Vd[2] + Vlin_gain +Vdrift_gain)/(V2[3]-V2[2])
SR104r 10k | alpha | 3.42e-07 | 1.9e-08 | 8 | SR104r_alpha |  | V1 | str_addr | COM2 | [6 empty] | where: | Vlin_gain is the linearity of DVMd
SR104r 10k | beta | -2.7e-08 | 5e-09 | 8 | SR104r_beta |  | V1 | role | switchbox | [7 empty] | Vdrift_gain is the zero-valued type-B drift of DVMd
SR104r 10k | gamma | 0 | 0 | inf | SR104r_gamma |  | V1 | test | A | [7 empty] | where: | uncertainty in Vdrift = abs(Vd[2]-(Vd[0]+((Vd[3]-Vd[2])/(V2[3]-V2[2]))*(V2[2]-V2[0])))/4
SR104r 10k | date | 2016-06-01 |  |  |  |  | Vd1 | addr | 2
SR104r 10k | T_sensor | SR104t 10k |  |  |  |  | Vd1 | str_addr | COM2
[7 empty] | Vd1 | role | switchbox
[7 empty] | Vd1 | test | C
[7 empty] | Vd2 | addr | 2
[7 empty] | Vd2 | str_addr | COM2
[7 empty] | Vd2 | role | switchbox
[7 empty] | Vd2 | test | B
[7 empty] | V2 | addr | 2
[7 empty] | V2 | str_addr | COM2
[7 empty] | V2 | role | switchbox
[7 empty] | V2 | test | D
[7 empty] | none | correction_100r | 0 | 0 | inf
[7 empty] | none | correction_10k | 0 | 0 | inf
[7 empty] | none | correction_100k | 0 | 0 | inf
[7 empty] | none | VRC_eq | 0 | 0 | inf
[7 empty] | none | VRC_10to1 | 0 | 0 | inf
[7 empty] | none | VRC_100to10 | 0 | 0 | inf
[7 empty] | none | linearity_gain | 0 | 0 | inf
[7 empty] | none | linearity_Vd | 0 | 0 | inf
[7 empty] | none | addr | 0
[7 empty] | none | str_addr | NO_ADDRESS
[7 empty] | none | role | NOT USED
[7 empty] | none | test
[7 empty] | DVM: HP34420A, s/n130 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | correction_10k | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | correction_100k | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | addr | 7
[7 empty] | DVM: HP34420A, s/n130 | str_addr | GPIB0::7::INSTR
[7 empty] | DVM: HP34420A, s/n130 | setfn_str | FUNC OHMF;OCOMP ON
[7 empty] | DVM: HP34420A, s/n130 | role | UNUSED HP34420A, s/n130
[7 empty] | DVM: HP34420A, s/n130 | test | *IDN?
[7 empty] | DVM: HP34401A, s/n976 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | correction_10k | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | correction_100k | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | addr | 17
[7 empty] | DVM: HP34401A, s/n976 | str_addr | GPIB0::17::INSTR
[7 empty] | DVM: HP34401A, s/n976 | setfn_str | FUNC OHMF;OCOMP ON
[7 empty] | DVM: HP34401A, s/n976 | role | UNUSED HP34401A, s/n976
[7 empty] | DVM: HP34401A, s/n976 | test | *IDN?
[7 empty] | DVM: HP3458A, s/n066 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP3458A, s/n066 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_066_10k_cor
[7 empty] | DVM: HP3458A, s/n066 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_066_100k_cor
[7 empty] | DVM: HP3458A, s/n066 | VRC_eq | 0 | 0 | 7 | VRCeq
[7 empty] | DVM: HP3458A, s/n066 | VRC_10to1 | 2.221e-07 | 9e-08 | 7 | VRC10to1 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n066 | VRC_100to10 | 4.809e-07 | 1.8e-07 | inf | VRC100to10 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n066 | linearity_gain | 0 | 2e-09 | 8 | Vlin_gain | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n066 | linearity_Vd | 0 | 2e-09 | 8 | Vlin_Vd | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n066 | addr | 0
[7 empty] | DVM: HP3458A, s/n066 | str_addr | GPIB0::0::INSTR
[7 empty] | DVM: HP3458A, s/n066 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n066 | role | UNUSED HP3458A, s/n066
[7 empty] | DVM: HP3458A, s/n066 | test | ID?
[7 empty] | DVM: HP3458A, s/n129 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP3458A, s/n129 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_129_10k_cor
[7 empty] | DVM: HP3458A, s/n129 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_066_100k_cor
[7 empty] | DVM: HP3458A, s/n129 | VRC_eq | 0 | 0 | 7 | VRCeq | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n129 | VRC_10to1 | 2.221e-07 | 9e-08 | 7 | VRC10to1 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n129 | VRC_100to10 | 4.809e-07 | 1.8e-07 | inf | VRC100to10 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n129 | linearity_gain | 0 | 2e-09 | 8 | Vlin_gain | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n129 | linearity_Vd | 0 | 2e-09 | 8 | Vlin_Vd
[7 empty] | DVM: HP3458A, s/n129 | addr | 25
[7 empty] | DVM: HP3458A, s/n129 | str_addr | GPIB0::25::INSTR
[7 empty] | DVM: HP3458A, s/n129 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n129 | role | UNUSED HP3458A, s/n129
[7 empty] | DVM: HP3458A, s/n129 | test | ID?
[7 empty] | DVM: HP3458A, s/n230 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP3458A, s/n230 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_230_10k_cor
[7 empty] | DVM: HP3458A, s/n230 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_230_100k_cor
[7 empty] | DVM: HP3458A, s/n230 | VRC_eq | 0 | 0 | 7 | VRCeq | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n230 | VRC_10to1 | 2.221e-07 | 9e-08 | 7 | VRC10to1 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n230 | VRC_100to10 | 4.809e-07 | 1.8e-07 | inf | VRC100to10 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n230 | linearity_gain | 0 | 2e-09 | 8 | Vlin_gain | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n230 | linearity_Vd | 0 | 2e-09 | 8 | Vlin_Vd
[7 empty] | DVM: HP3458A, s/n230 | addr | 22
[7 empty] | DVM: HP3458A, s/n230 | str_addr | GPIB0::22::INSTR
[7 empty] | DVM: HP3458A, s/n230 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n230 | role | UNUSED HP3458A, s/n230
[7 empty] | DVM: HP3458A, s/n230 | test | ID?
[7 empty] | DVM: HP3458A, s/n452 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP3458A, s/n452 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_452_10k_cor
[7 empty] | DVM: HP3458A, s/n452 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_452_100k_cor
[7 empty] | DVM: HP3458A, s/n452 | VRC_eq | 0 | 0 | 7 | VRCeq | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n452 | VRC_10to1 | 2.221e-07 | 9e-08 | 7 | VRC10to1 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n452 | VRC_100to10 | 4.809e-07 | 1.8e-07 | inf | VRC100to10 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n452 | linearity_gain | 0 | 2e-09 | 8 | Vlin_gain | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n452 | linearity_Vd | 0 | 2e-09 | 8 | Vlin_Vd
[7 empty] | DVM: HP3458A, s/n452 | addr | 23
[7 empty] | DVM: HP3458A, s/n452 | str_addr | GPIB0::23::INSTR
[7 empty] | DVM: HP3458A, s/n452 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n452 | role | UNUSED HP3458A, s/n452
[7 empty] | DVM: HP3458A, s/n452 | test | ID?
[7 empty] | DVM: HP3458A, s/n518 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP3458A, s/n518 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_518_10k_cor
[7 empty] | DVM: HP3458A, s/n518 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_518_100k_cor
[7 empty] | DVM: HP3458A, s/n518 | VRC_eq | 0 | 0 | 7 | VRCeq | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n518 | VRC_10to1 | 2.221e-07 | 9e-08 | 7 | VRC10to1 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n518 | VRC_100to10 | 4.809e-07 | 1.8e-07 | inf | VRC100to10 | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n518 | linearity_gain | 0 | 2e-09 | 8 | Vlin_gain | (link to RefStep data?)
[7 empty] | DVM: HP3458A, s/n518 | linearity_Vd | 0 | 2e-09 | 8 | Vlin_Vd
[7 empty] | DVM: HP3458A, s/n518 | addr | 24
[7 empty] | DVM: HP3458A, s/n518 | str_addr | GPIB0::24::INSTR
[7 empty] | DVM: HP3458A, s/n518 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n518 | role | UNUSED HP3458A, s/n518
[7 empty] | DVM: HP3458A, s/n518 | test | ID?
[7 empty] | SRC: D4808 | addr | 2
[7 empty] | SRC: D4808 | str_addr | GPIB0::2::INSTR
[7 empty] | SRC: D4808 | init_str | F0G0D0S0= | M+0O1= | R0=
[7 empty] | SRC: D4808 | setV_str | M | =
[7 empty] | SRC: D4808 | oper_str | O1=
[7 empty] | SRC: D4808 | stby_str | O0=
[7 empty] | SRC: D4808 | role | UNUSED D4808
[7 empty] | SRC: D4808 | test | X8=
[7 empty] | SRC: F5520A | addr | 4
[7 empty] | SRC: F5520A | str_addr | GPIB0::4::INSTR
[7 empty] | SRC: F5520A | init_str | OUT 0V,OPER
[7 empty] | SRC: F5520A | setV_str | OUT | V,0Hz
[7 empty] | SRC: F5520A | oper_str | OPER
[7 empty] | SRC: F5520A | stby_str | STBY
[7 empty] | SRC: F5520A | chk_err_str | ERR? | *CLS
[7 empty] | SRC: F5520A | role | UNUSED F5520A
[7 empty] | SRC: F5520A | test | *IDN?
Note on F2: Tim Lawson: As the resistor section grows and is added to by HRBA, the label will be appended with the RunId of the run that was the source of these values. The combined label acts as the AnalysisId.
Note on B5: Tim Lawson: IUnits of [deg C]^-1
Note on B8: Tim Lawson: This means any permanently-attached temperature sensor that forms part of the resistance standard. In practice this is 'none' for any RTD (eg Pt 100r, SR104t) and units that have a thermometer well but no permanant sensor (eg:Tinsleys). Typically 'Pt100r' is used for 'home-made' resistance boxes. In future, each Pt will need a unique id (eg: 'H100M_Pt 100r', 'New10M_Pt 100r', etc.); calibration info and a prediction equation in the same manner of SR104t 10k.
Note on A15: Tim Lawson: From Buildup2015.xls Since this has a full years data and higher dof

## Results
Start row | 4