## Comments
WARNING: DO NOT INSERT ANY COMMENTS IN THIS WORKBOOK!
Due to a possible bug in the Python module Openpyxl, used to read & write .xlsx workbooks, the presence of any comment causes file corruption on re-saving.
Below are the text of all comments that have now been removed.
Cells that have had comments removed are now have formatted with a thick border.
Sheet | Cell | Text
Data | A1 | Tim Lawson: Make sure these are updated before running HRBC or HRBA! Otherwise the wrong settings will be run or analysed.
Data | A4 | Tim Lawson: The Run Id identifies a particular physical setup and its only use is to allow the analysis program (HRBA) to match measurement data with corresponding Rlink data.
Data | A4 |  | Each Run Id must be unique and is automatically generated by combining run information such as: the version of HRBC used to acquire the data; the names of the two resistors used and the date and time when the run was started.
Data | A4 |  | The Run Id should NOT be confused with the comment, which is intended to record and describe (as opposed to identify) the physical setup, along with any other relevant observations.
Data | A4 |  | The comment is intended to be read by a human, not HRBA.
Rlink | A1 | Tim Lawson: 1st row of actual data (after 6 lines of header)
Rlink | C1 | Tim Lawson: Make sure these are updated to reflect nominal resister ratio.
Rlink | C2 | Tim Lawson: Make sure these are updated to reflect nominal resister ratio.
Parameters | F2 | Tim Lawson: As the resistor section grows and is added to by HRBA, the label will be appended with the RunId of the run that was the source of the values. The combined label acts as the Analysis Id.
Parameters | B5 | Tim Lawson: Units of [deg C]^-1
Parameters | B8 | Tim Lawson: This means any permanently-attached temperature sensor that forms part of the resistance standard.
Parameters | B8 |  | In practice this is 'none' for any RTD (eg Pt 100r, SR104t) and units that have a thermometer well but no permanant sensor (eg:Tinsleys).
Parameters | B8 |  | Typically 'Pt100r' is used for 'home-made' resistance boxes. In future, each Pt will need a unique id (eg: 'H100M_Pt 100r', 'New10M_Pt 100r', etc.); calibration info and a prediction equation in the same manner of SR104t 10k.
Parameters | A15 | Tim Lawson: Buildup2015.xls - Since this has a full years data and higher dof.
Parameters | K72 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K73 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K87 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K88 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K102 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K103 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K117 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K118 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K132 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Parameters | K133 | Tim Lawson: = [Max Linearity of 10V range at +/- 1V] x [range ratio], where range ratio is 100mV/10V. Units are [V].
Results | B1 | Tim Lawson: Two rows above 1st row of actual data.

## Data
start_row | 6
Run Id: | HRBC.v1.0 G493 100k:SR104r 10k 03/03/2017 11:37:33 | [5 empty] | V1 |  | V2 |  | Va |  | Vb |  | Vc |  | Vd |  | dvm_T1 | dvm_T2 | GMH_T1 | GMH_T2 | Ambient conditions |  |  | Time | Comment | Role | Instrument descr.
V1_set | V2_set | n | Start/xl del. | AZ1 del. | Range del. | Relay del. | V_meas | sd(V) | V_meas | sd(V) | V_meas | sd(V) | V_meas | sd(V) | V_meas | sd(V) | V_meas | sd(V) | (Ohm) | (Ohm) | (deg C) | (deg C) | T (°C) | P(mbar) | RH(%)
10 | -1.0004 | 10 | 60 | 5 | 90 | 5 | [21 empty] | SRC2 | SRC: F5520A
-10 | 1.0004 | 10 | 60 | 5 | 90 | 5 | [21 empty] | DVMT1 | none
10 | -1.0004 | 10 | 60 | 5 | 90 | 5 | [21 empty] | SRC1 | SRC: D4808
10 | -1.0014 | 10 | 60 | 5 | 90 | 5 | [21 empty] | GMH1 | GMH: s/n628
-10 | 1.0014 | 10 | 60 | 5 | 90 | 5 | [21 empty] | GMH2 | GMH: s/n627
10 | -1.0014 | 10 | 60 | 5 | 90 | 5 | [21 empty] | DVMT2 | DVM: HP34401A, s/n976
[28 empty] | DVM | DVM: HP3458A, s/n452
[28 empty] | switchbox | V1
[28 empty] | GMHroom | GMH: s/n367

## Rlink
Start row | 28 | ¦V1_set¦ | 10
N_Reversals | 10 | ¦V2_set¦ | 1
N_Readings | 10
Run Id: | HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33
Comment | R1: Ti612 100k monitored by GMH, s/n628. R2: SR104r 10k monitored by GMH, s/n627. 1st 'proper' run after some bug-fixing.
06/03/2017 10:33:16 |  | Nom. value | ¦V¦
R1 | Ti612 100k | 100000 | 10
R2 | SR104r 10k | 10000 | 1
ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
-2.5508e-07 | -1.3577e-06 | 7.7893e-08 | -1.5265e-06 | -9.0641e-08 | -8.5265e-07 | -4.5403e-07 | -1.1512e-06 | -9.3348e-07 | -2.5551e-07
-2.3298e-07 | -8.848e-07 | 4.6963e-07 | -1.1438e-06 | 2.3672e-07 | -9.0812e-07 | -4.6877e-07 | -1.0724e-06 | 1.5667e-07 | -1.1489e-06
3.1379e-08 | -8.8662e-07 | -3.2811e-07 | -1.0835e-06 | -2.7552e-07 | -1.8879e-06 | 6.1374e-07 | -1.5178e-06 | 3.3649e-07 | -1.241e-06
-3.4041e-07 | -4.5753e-07 | 3.3335e-07 | -6.7702e-07 | -1.4015e-07 | -1.1723e-06 | -1.8197e-07 | -1.0005e-06 | 2.7491e-07 | -5.7806e-07
1.3132e-07 | -3.1457e-07 | 1.8785e-07 | -1.1138e-06 | -4.662e-07 | -1.4292e-06 | 3.2799e-07 | -1.0956e-06 | -5.4079e-07 | -5.7905e-07
2.6953e-07 | -1.2124e-06 | -7.9743e-07 | -3.648e-07 | -3.0417e-07 | -6.0672e-07 | 3.2218e-07 | -1.2665e-06 | 4.6933e-07 | -7.6162e-07
-1.576e-08 | -2.8323e-07 | -2.1856e-07 | -1.663e-06 | 6.065e-07 | -1.0702e-06 | 1.9884e-07 | -6.167e-07 | -7.8463e-07 | -1.3153e-06
-1.2309e-07 | -1.2379e-06 | 6.4553e-07 | -2.3576e-06 | 6.7141e-08 | -8.065e-07 | -4.3257e-08 | -1.2472e-06 | 2.63e-07 | -1.6592e-06
-2.0469e-08 | -1.2239e-06 | -1.1623e-08 | -1.6839e-06 | -2.5092e-07 | -1.2076e-06 | -3.2335e-07 | -5.5927e-07 | 8.5402e-09 | -1.0974e-06
5.0666e-07 | -9.176e-07 | 7.4889e-07 | -1.298e-06 | -3.686e-07 | -1.0863e-06 | 1.6639e-08 | -1.2377e-06 | -5.4359e-07 | -2.6676e-07
Run Id: | TestPoint data for HRBC validation
Comment | R1: Ti611 1M monitored by GMH: s/n628. R2: Ti612 100k monitored by GMH: s/n627
2017-08-04 |  | Nom. value | ¦V¦
R1 | Ti611 1M | 1000000 | 20
R2 | Ti612 100k | 100000 | 2
ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
7.7869e-08 | 3.2664e-07 | -1.2537e-06 | -3.0338e-07 | -5.8365e-07 | -1.8821e-06 | -2.6724e-06 | -1.6426e-06 | -9.3634e-07 | -2.3927e-06
4.5632e-08 | 2.8439e-07 | -1.2693e-06 | -2.9749e-07 | -5.7541e-07 | -1.8799e-06 | -2.6762e-06 | -1.5743e-06 | -1.0541e-06 | -2.3957e-06
-6.8742e-08 | 2.4435e-07 | -1.2879e-06 | -2.8336e-07 | -5.495e-07 | -1.8976e-06 | -2.8234e-06 | -1.5537e-06 | -9.6666e-07 | -2.4101e-06
-9.9655e-08 | 3.0338e-07 | -1.317e-06 | -3.0883e-07 | -5.732e-07 | -1.9302e-06 | -2.7716e-06 | -1.6058e-06 | -8.8629e-07 | -2.411e-06
-8.1402e-08 | 2.8189e-07 | -1.3641e-06 | -2.0652e-07 | -6.3635e-07 | -1.9741e-06 | -2.8025e-06 | -1.6511e-06 | -9.2309e-07 | -2.4238e-06
-1.575e-08 | 2.9057e-07 | -1.4061e-06 | -1.6501e-07 | -6.7756e-07 | -1.9834e-06 | -2.8078e-06 | -1.6176e-06 | -1.1389e-06 | -2.4287e-06
1.2365e-08 | 3.1e-07 | -1.4833e-06 | -2.1535e-07 | -7.2319e-07 | -1.9679e-06 | -2.7884e-06 | -1.6117e-06 | -1.0978e-06 | -2.3963e-06
-1.5059e-07 | 3.5608e-07 | -1.4339e-06 | -1.6766e-07 | -6.7079e-07 | -1.9544e-06 | -2.7001e-06 | -1.6167e-06 | -1.0545e-06 | -2.3506e-06
-1.185e-07 | 3.4003e-08 | -1.4233e-06 | -1.263e-07 | -6.7226e-07 | -1.9314e-06 | -2.7734e-06 | -1.5923e-06 | -9.8933e-07 | -2.4708e-06
-2.0358e-07 | 2.473e-08 | -1.4624e-06 | -8.4346e-08 | -6.9935e-07 | -1.9859e-06 | -2.7752e-06 | -1.6438e-06 | -9.1426e-07 | -2.4578e-06
Run Id: | HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48
Comment | R1: Ti611 1M monitored by GMH: s/n628. R2: Ti612 100k monitored by GMH: s/n627
03/08/2017 10:51:09 |  | Nom. value | ¦V¦
R1 | Ti611 1M | 1000000 | 10
R2 | Ti612 100k | 100000 | 1
ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
4.2713e-08 | -3.0505e-07 | 2.2404e-08 | -2.7803e-07 | -4.0051e-07 | -1.3091e-07 | -3.1831e-07 | -7.8235e-07 | -2.7351e-07 | -3.2007e-07
-9.178e-08 | -4.3968e-07 | 4.805e-08 | -3.1288e-07 | -2.1497e-08 | -2.3076e-07 | -1.0966e-07 | -2.469e-07 | -2.093e-07 | -3.9223e-07
-1.3069e-07 | -2.9034e-07 | -1.4298e-08 | -2.6998e-07 | -3.4506e-07 | -2.6291e-07 | -1.1291e-07 | -2.1228e-07 | -3.6781e-07 | -1.5241e-07
-9.413e-08 | -2.5925e-07 | -2.5601e-07 | -2.6884e-07 | -2.1651e-07 | -2.1139e-07 | -6.3537e-08 | -2.7246e-07 | -1.8809e-07 | -1.5635e-07
-6.0519e-08 | -3.56e-07 | -2.505e-08 | -1.6167e-07 | -7.9062e-08 | -4.2364e-07 | -2.2862e-07 | -3.5734e-07 | -4.0638e-08 | -2.2794e-07
-1.3055e-07 | 6.9948e-08 | -2.7793e-07 | -1.0117e-07 | -3.7851e-08 | -2.7457e-07 | -3.1824e-07 | -3.3148e-07 | -2.1524e-07 | -3.0802e-07
-1.5702e-07 | -2.3263e-07 | -2.6204e-08 | -1.2452e-07 | -1.2359e-07 | -4.2494e-07 | -1.3339e-07 | -2.8632e-07 | -1.8226e-07 | -5.4631e-07
-1.9572e-07 | -1.9434e-07 | -3.1587e-07 | -2.8986e-07 | -1.6709e-07 | -4.1106e-07 | -2.4195e-07 | 6.2042e-08 | -5.9424e-07 | -4.3647e-07
-7.1965e-08 | -2.1176e-07 | -2.4229e-07 | -5.5445e-07 | -1.0271e-07 | -2.702e-07 | -2.2199e-07 | -3.3362e-07 | -3.0749e-07 | -4.8075e-07
-3.0763e-07 | -2.1148e-07 | -7.6404e-08 | -4.1864e-07 | -1.0919e-07 | -4.5911e-07 | -1.7385e-07 | -2.7645e-07 | -8.3079e-08 | -2.8349e-07
[10 empty] | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-
[10 empty] | ΔV+ | ΔV- | ΔV+ | ΔV- | ΔV+ | ΔV-

## Parameters
Resistor Info: | [6 empty] | Instrument Info: | [6 empty] | Notes:
description | parameter | value | uncert | dof | label | Comment / Reference | description | parameter | value | uncert | dof | label | Comment / Reference | 1. Units of alpha and beta are [deg C]^-1 and [def C]^-2, respectively (i.e. 1e-6*ppm/C...).
Pt 100r | R0_LV | 100.0118 | 0 | inf | Pt_R0 | Don't use for absolute temperature measurement - differences only! | GMH: s/n367 | T_correction |  |  |  | GMH367T_cor | (Cal report?) | 2. Units of gamma are [V]^-1 (i.e. 1e-6*ppm/V).
Pt 100r | TRef_LV | 0 | 0 | inf | Pt_Tref_LV |  | GMH: s/n367 | P_correction |  |  |  | GMH367P_cor | (Cal report?) | 3. Uncertainties in Rs propagate through to dependant Rx values through the build-up. This is evident in the increasing dof as Rx increases.
Pt 100r | alpha | 0.0039091 | 0 | inf | Pt_alpha |  | GMH: s/n367 | RH_correction |  |  |  | GMH367RH_cor | (Cal report?) | 4. For the Hamon, the voltage coefficient is the effect as it [is/would be] measured on the complete configuration of elements;
Pt 100r | beta | -5.926e-07 | 0 | inf | Pt_beta |  | GMH: s/n367 | addr | 8 |  |  |  | Windows-assigned port to 3100N USB connector | The series configuration has 1/10 the voltage drop per element, compared with the parallel, for the same overall ppm change in resistance.
Pt 100r | date | 1899-12-31 |  |  |  |  | GMH: s/n367 | str_addr | COM8 |  |  |  | Windows-assigned port to 3100N USB connector | 5. Calibration dates relate to the LV measurement.
Pt 100r | T_sensor | none |  |  |  |  | GMH: s/n367 | role | UNUSED GMH, s/n367 |  |  |  |  | This shouldn't cause a problem as the LV and HV meaurements are interleaved and their means are not likely to differ by more than a few minutes.
SR104t 10k | R0_LV | 10000.2 | 0 | inf | SR104t_R0 |  | GMH: s/n367 | test | RH |  |  |  |  | 6. Measurement equation(s):
SR104t 10k | TRef_LV | 23 | 0 | inf | SR104t_Tref_LV |  | GMH: s/n627 | T_correction | 0.01 | 0.0086 | 18.1 | GMH627_cor | Temperature/2018/723, units are deg C |  | R1 = -(R2*(1+vrc)*V1_av*G)/( G*V2_av -Vd_av)
SR104t 10k | alpha | 0.001 | 0 | inf | SR104t_alpha |  | GMH: s/n627 | addr | 9 |  |  |  | Windows-assigned port to 3100N USB connector |  | where: | R2 = R2_0*(1+R2alpha*dT2 + R2beta*dT2**2 + R2gamma*dV2) + Rd
SR104t 10k | beta | 0 | 0 | inf | SR104t_beta |  | GMH: s/n627 | str_addr | COM9 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | where: | dT2 = T2_av - R2TRef + T_def2
SR104t 10k | date | 1899-12-31 |  |  |  |  | GMH: s/n627 | role | UNUSED GMH, s/n627 | [7 empty] | ( where: | T2_av = T2_av_gmh )
SR104t 10k | T_sensor | none |  |  |  |  | GMH: s/n627 | test | T |  |  |  | Test Temperature measurement function |  |  |  | dV2 = abs(abs(V2av) - R2VRef)
SR104r 10k | R0_LV | 9999.9993 | 0.00033 | 16 | SR104r_R0_LV <source?> |  | GMH: s/n628 | T_correction | 0 | 0.0076 | 18.1 | GMH628_cor | Temperature/2018/724, units are deg C |  |  | V1av = (V1[0]-2*V1[1]+V1[2])/4
SR104r 10k | TRef_LV | 20 | 0 | inf | SR104r_Tref_LV <source?> |  | GMH: s/n628 | addr | 5 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | V2av = (V2[0]-2*V2[1]+V2[2])/4
SR104r 10k | VRef_LV | 5 | 0 | inf | SR104r_Vref_LV <source?> |  | GMH: s/n628 | str_addr | COM5 |  |  |  | Windows-assigned port to 3100N USB connector |  |  | Vdav = (Vd[0]-2*Vd[1]+Vd[2])/4 + Vlin_Vd + Vdrift_Vd
SR104r 10k | R0_HV | 9999.9993 | 0.00033 | 16 | SR104r_R0_HV <source?> |  | GMH: s/n628 | role | UNUSED GMH, s/n628 | [6 empty] | where: | Vlin_Vd is the linearity of DVMd
SR104r 10k | TRef_HV | 20 | 0 | inf | SR104r_Tref_HV <source?> |  | GMH: s/n628 | test | T |  |  |  | Test Temperature measurement function |  |  |  | Vdrift_Vd is the zero-valued type-B drift of DVMd
SR104r 10k | VRef_HV | 10 | 0 | inf | SR104r_Vref_HV <source?> |  | V1 | addr | 2 | [6 empty] | G = (Vd[3]-Vd[2] + Vlin_gain +Vdrift_gain)/(V2[3]-V2[2])
SR104r 10k | alpha | 3.42e-07 | 1.9e-08 | 8 | SR104r_alpha <source?> |  | V1 | str_addr | COM2 | [6 empty] | where: | Vlin_gain is the linearity of DVMd
SR104r 10k | beta | -2.7e-08 | 5e-09 | 8 | SR104r_beta <source?> |  | V1 | role | switchbox | [7 empty] | Vdrift_gain is the zero-valued type-B drift of DVMd
SR104r 10k | gamma | 0 | 0 | inf | SR104r_gamma <source?> |  | V1 | test | A | [7 empty] | where: | uncertainty in Vdrift = abs(Vd[2]-(Vd[0]+((Vd[3]-Vd[2])/(V2[3]-V2[2]))*(V2[2]-V2[0])))/4
SR104r 10k | date | 2016-06-01 |  |  |  |  | Vd1 | addr | 2
SR104r 10k | T_sensor | SR104t 10k |  |  |  |  | Vd1 | str_addr | COM2
Ti612 100k | R0_LV | 100003.6255 | 0.21111 | 77 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd1 | role | switchbox
Ti612 100k | TRef_LV | 20.7904 | 0.0014778 | 13 | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd1 | test | C
Ti612 100k | VRef_LV | 10 | 9.1808e-08 | 147 | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd2 | addr | 2
Ti612 100k | R0_HV | 100003.6255 | 0.21111 | 77 | Ti612_R0_HV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd2 | str_addr | COM2
Ti612 100k | TRef_HV | 20.7904 | 0.0014778 | 13 | Ti612_TRef_HV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd2 | role | switchbox
Ti612 100k | VRef_HV | 10 | 9.1808e-08 | 147 | Ti612_VRef_HV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | Vd2 | test | B
Ti612 100k | alpha | 8.95e-06 | 1.7e-07 | 8.4 | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | …\E052\Proc Devel\CCEM-K2(2012)_TBL.xlsx | V2 | addr | 2
Ti612 100k | beta | 0 | 0 | inf | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | V2 | str_addr | COM2
Ti612 100k | gamma | 0 | 0 | inf | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | HRBA1.2_13/06/2017 11:21:25 | V2 | role | switchbox
Ti612 100k | date | 2017-03-03 13:43 |  |  |  | HRBA1.2_13/06/2017 11:21:25 | V2 | test | D
Ti612 100k | T_sensor | none |  |  |  | HRBA1.2_13/06/2017 11:21:25 | none | correction_100r | 0 | 0 | inf
[7 empty] | none | correction_10k | 0 | 0 | inf
[7 empty] | none | correction_100k | 0 | 0 | inf
[7 empty] | none | VRC_eq | 0 | 0 | inf
[7 empty] | none | VRC_10to1 | 0 | 0 | inf
[7 empty] | none | VRC_100to10 | 0 | 0 | inf
[7 empty] | none | linearity_gain | 0 | 0 | inf
[7 empty] | none | linearity_Vd | 0 | 0 | inf
[7 empty] | none | addr | 0
[7 empty] | none | str_addr | NO_ADDRESS
[7 empty] | none | role | NOT USED
[7 empty] | none | test
[7 empty] | DVM: HP34420A, s/n130 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | correction_10k | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | correction_100k | 0 | 0 | inf
[7 empty] | DVM: HP34420A, s/n130 | addr | 7
[7 empty] | DVM: HP34420A, s/n130 | str_addr | GPIB0::7::INSTR
[7 empty] | DVM: HP34420A, s/n130 | setfn_str | FUNC OHMF;OCOMP ON
[7 empty] | DVM: HP34420A, s/n130 | role | UNUSED HP34420A, s/n130
[7 empty] | DVM: HP34420A, s/n130 | test | *IDN?
[7 empty] | DVM: HP34401A, s/n976 | correction_100r | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | correction_10k | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | correction_100k | 0 | 0 | inf
[7 empty] | DVM: HP34401A, s/n976 | addr | 17
[7 empty] | DVM: HP34401A, s/n976 | str_addr | GPIB0::17::INSTR
[7 empty] | DVM: HP34401A, s/n976 | setfn_str | FUNC OHMF;OCOMP ON
[7 empty] | DVM: HP34401A, s/n976 | role | UNUSED HP34401A, s/n976
[7 empty] | DVM: HP34401A, s/n976 | test | *IDN?
[7 empty] | DVM: HP3458A, s/n066 | correction_100r | 0 | 0 | inf | 3458A_066_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_066_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_066_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | correction_100k | 0 | 0 | inf | 3458A_066_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | Vgain_05r10 | 1 | 1.4142e-06 | 120.8751 | 3458A_066_Vgain_01r01 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | Vgain_05r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_066_Vgain_05r1 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_066_Vgain_1r1 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | Vgain_1r10 | 1 | 8.1317e-07 | 120.8751 | 3458A_066_Vgain_1r10 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | Vgain_5r10 | 1 | 3.5355e-07 | 120.8751 | 3458A_066_Vgain_5r10 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_066_Vgain_10r10 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | 3458A_066_Vgain_50r100 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n066 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_066_Vgain_100r100 | Vgain_calcs.xlsx, S21939
[7 empty] | DVM: HP3458A, s/n066 | linearity_pert | 0 | 2.9e-08 | 18 | 3458A_066_Vlin_gain | S21939
[7 empty] | DVM: HP3458A, s/n066 | linearity_Vdav | 0 | 2.9e-08 | 18 | 3458A_066_Vlin_Vd | S21939
[7 empty] | DVM: HP3458A, s/n066 | addr | 0
[7 empty] | DVM: HP3458A, s/n066 | str_addr | GPIB0::0::INSTR
[7 empty] | DVM: HP3458A, s/n066 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n066 | role | UNUSED HP3458A, s/n066
[7 empty] | DVM: HP3458A, s/n066 | test | ID?
[7 empty] | DVM: HP3458A, s/n129 | correction_100r | 0 | 0 | inf | 3458A_129_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n129 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_129_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n129 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_129_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n129 | Vgain_05r10 | 1 | 1.4142e-06 | 120.8751 | 3458A_129_Vgain_01r01 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n129 | Vgain_05r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_129_Vgain_05r1 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_129_Vgain_1r1 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | Vgain_1r10 | 1 | 8.1317e-07 | 120.8751 | 3458A_129_Vgain_1r10 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | Vgain_5r10 | 1 | 3.5355e-07 | 120.8751 | 3458A_129_Vgain_5r10 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_129_Vgain_10r10 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | 3458A_129_Vgain_50r100 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n129 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_129_Vgain_100r100 | Vgain_calcs.xlsx, S21930
[7 empty] | DVM: HP3458A, s/n129 | linearity_pert | 0 | 2e-08 | 18 | 3458A_129_Vlin_gain | S21930
[7 empty] | DVM: HP3458A, s/n129 | linearity_Vdav | 0 | 2e-08 | 18 | 3458A_129_Vlin_Vd | S21930
[7 empty] | DVM: HP3458A, s/n129 | addr | 25
[7 empty] | DVM: HP3458A, s/n129 | str_addr | GPIB0::25::INSTR
[7 empty] | DVM: HP3458A, s/n129 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n129 | role | UNUSED HP3458A, s/n129
[7 empty] | DVM: HP3458A, s/n129 | test | ID?
[7 empty] | DVM: HP3458A, s/n230 | correction_100r | 0 | 0 | inf | 3458A_230_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n230 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_230_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n230 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_230_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n230 | Vgain_05r10 | 1 | 1.4142e-06 | 120.8751 | 3458A_230_Vgain_01r01 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n230 | Vgain_05r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_230_Vgain_05r1 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_230_Vgain_1r1 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | Vgain_1r10 | 1 | 8.1318e-07 | 120.8751 | 3458A_230_Vgain_1r10 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | Vgain_5r10 | 1 | 3.5356e-07 | 120.8751 | 3458A_230_Vgain_5r10 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_230_Vgain_10r10 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | 3458A_230_Vgain_10r100 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n230 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_230_Vgain_100r100 | Vgain_calcs.xlsx, S21928
[7 empty] | DVM: HP3458A, s/n230 | linearity_pert | 0 | 2e-08 | 18 | 3458A_230_Vlin_gain | S21928
[7 empty] | DVM: HP3458A, s/n230 | linearity_Vdav | 0 | 2e-08 | 18 | 3458A_230_Vlin_Vd | S21928
[7 empty] | DVM: HP3458A, s/n230 | addr | 22
[7 empty] | DVM: HP3458A, s/n230 | str_addr | GPIB0::22::INSTR
[7 empty] | DVM: HP3458A, s/n230 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n230 | role | UNUSED HP3458A, s/n230
[7 empty] | DVM: HP3458A, s/n230 | test | ID?
[7 empty] | DVM: HP3458A, s/n382 | correction_100r | 0 | 0 | inf | 3458A_382_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n382 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_382_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n382 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_382_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n382 | Vgain_0.1r0.1 | 1 | 1.4142e-06 | 120.8751 | 3458A_382_Vgain_0.1r0.1 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_0.5r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_382_Vgain_0.5r1 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_382_Vgain_1r1 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_1r10 | 1 | 7.0711e-07 | 120.8751 | 3458A_382_Vgain_1r10 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_5r10 | 1 | 3.5355e-07 | 120.8751 | 3458A_382_Vgain_5r10 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_382_Vgain_10r10 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | *3458A_382_Vgain_10r100* | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n382 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_382_Vgain_100r100 | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | linearity_pert | 0 | 1.4e-08 | 18 | 3458A_382_Vlin_gain | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | linearity_Vdav | 0 | 1.4e-08 | 18 | 3458A_382_Vlin_Vd | Vgain_calcs.xlsx
[7 empty] | DVM: HP3458A, s/n382 | addr | 22
[7 empty] | DVM: HP3458A, s/n382 | str_addr | GPIB0::22::INSTR
[7 empty] | DVM: HP3458A, s/n382 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n382 | role | UNUSED HP3458A, s/n382
[7 empty] | DVM: HP3458A, s/n382 | test | ID?
[7 empty] | DVM: HP3458A, s/n452 | correction_100r | 0 | 0 | inf | 3458A_452_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n452 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_452_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n452 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_452_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n452 | Vgain_05r10 | 1 | 1.4142e-06 | 120.8751 | 3458A_452_Vgain_01r01 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n452 | Vgain_05r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_452_Vgain_05r1 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n452 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_452_Vgain_1r1 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n452 | Vgain_1r10 | 1 | 8.1318e-07 | 120.8751 | 3458A_452_Vgain_1r10 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n452 | Vgain_5r10 | 1 | 3.5355e-07 | 120.8751 | 3458A_452_Vgain_5r10 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n452 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_452_Vgain_10r10 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n518 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | 3458A_452_Vgain_10r100 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_452_Vgain_100r100 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n452 | linearity_pert | 0 | 1.9e-08 | 18 | 3458A_452_Vlin_gain | S21927
[7 empty] | DVM: HP3458A, s/n452 | linearity_Vdav | 0 | 1.9e-08 | 18 | 3458A_452_Vlin_Vd | S21927
[7 empty] | DVM: HP3458A, s/n452 | addr | 23
[7 empty] | DVM: HP3458A, s/n452 | str_addr | GPIB0::23::INSTR
[7 empty] | DVM: HP3458A, s/n452 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n452 | role | UNUSED HP3458A, s/n452
[7 empty] | DVM: HP3458A, s/n452 | test | ID?
[7 empty] | DVM: HP3458A, s/n518 | correction_100r | 0 | 0 | inf | 3458A_518_100r_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | correction_10k | -5.68e-05 | 2e-07 | inf | 3458A_518_10k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | correction_100k | -5.63e-05 | 6e-07 | inf | 3458A_518_100k_cor | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | Vgain_05r10 | 1 | 1.4142e-06 | 120.8751 | 3458A_518_Vgain_05r10 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | Vgain_05r1 | 1 | 9.1924e-07 | 120.8751 | 3458A_518_Vgain_05r1 | Vgain_calcs.xlsx, S21926
[7 empty] | DVM: HP3458A, s/n518 | Vgain_1r1 | 1 | 7.0711e-07 | 120.8751 | 3458A_518_Vgain_1r1 | Vgain_calcs.xlsx, S21927
[7 empty] | DVM: HP3458A, s/n518 | Vgain_1r10 | 1 | 8.1319e-07 | 120.8751 | 3458A_518_Vgain_1r10 | Vgain_calcs.xlsx, S21926
[7 empty] | DVM: HP3458A, s/n518 | Vgain_5r10 | 1 | 3.5355e-07 | 120.8751 | 3458A_518_Vgain_5r10 | Vgain_calcs.xlsx, S21926
[7 empty] | DVM: HP3458A, s/n518 | Vgain_10r10 | 1 | 2.4749e-07 | 120.8751 | 3458A_518_Vgain_10r10 | Vgain_calcs.xlsx, S21926
[7 empty] | DVM: HP3458A, s/n518 | Vgain_10r100 | 1 | 8.4853e-07 | 120.8751 | 3458A_518_Vgain_10r100 | UNKNOWN!
[7 empty] | DVM: HP3458A, s/n518 | Vgain_100r100 | 1 | 7.0711e-07 | 120.8751 | 3458A_518_Vgain_100r100 | Vgain_calcs.xlsx, S21926
[7 empty] | DVM: HP3458A, s/n518 | linearity_pert | 0 | 1.7e-08 | 18 | 3458A_518_Vlin_gain | S21926
[7 empty] | DVM: HP3458A, s/n518 | linearity_Vdav | 0 | 1.7e-08 | 18 | 3458A_518_Vlin_Vd | S21926
[7 empty] | DVM: HP3458A, s/n518 | addr | 24
[7 empty] | DVM: HP3458A, s/n518 | str_addr | GPIB0::24::INSTR
[7 empty] | DVM: HP3458A, s/n518 | setfn_str | DCV
[7 empty] | DVM: HP3458A, s/n518 | role | UNUSED HP3458A, s/n518
[7 empty] | DVM: HP3458A, s/n518 | test | ID?
[7 empty] | SRC: D4808 | addr | 2
[7 empty] | SRC: D4808 | str_addr | GPIB0::2::INSTR
[7 empty] | SRC: D4808 | init_str | F0G0D0S0= | M+0O1= | R0=
[7 empty] | SRC: D4808 | setV_str | M | =
[7 empty] | SRC: D4808 | oper_str | O1=
[7 empty] | SRC: D4808 | stby_str | O0=
[7 empty] | SRC: D4808 | role | UNUSED D4808
[7 empty] | SRC: D4808 | test | X8=
[7 empty] | SRC: F5520A | addr | 4
[7 empty] | SRC: F5520A | str_addr | GPIB0::4::INSTR
[7 empty] | SRC: F5520A | init_str | OUT 0V,OPER
[7 empty] | SRC: F5520A | setV_str | OUT | V,0Hz
[7 empty] | SRC: F5520A | oper_str | OPER
[7 empty] | SRC: F5520A | stby_str | STBY
[7 empty] | SRC: F5520A | chk_err_str | ERR? | *CLS
[7 empty] | SRC: F5520A | role | UNUSED F5520A
[7 empty] | SRC: F5520A | test | *IDN?

## Results
Start row | 409
Processed with HRBA v1.2 on Tuesday, 13. June 2017 11:21AM
|  | Run Id: | HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 |  |  |  |  | Uncertainty Budget | [6 empty] | R1(T) |  |  | exp. U(95%) | av T |  |  | av date/time | av V
Name | Test V | Date | T | R1 | std u. | dof | exp. U(95%) | Quantity (Label) | Value | Std u | dof | sens. coef. | u contrib. | LV | 100003.6255 | 0.21111 | 77 | 0.42041 | 20.7904 | 0.0014778 | 13 | 03/03/2017 13:42:59 | 10 | 9.1808e-08 | 147
Ti612 100k | 10 | 03/03/2017 11:52:42 | 20.7646 | 100003.5689 | 0.21885 | 87 | 0.43498 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.5959 | 0.21001 | HV | 100003.6255 | 0.21111 | 77 | 0.42041 | 20.7904 | 0.0014778 | 13 | 03/03/2017 13:42:59 | 10 | 9.1808e-08 | 147
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 7.3424e-07 | 9 | -50003.4369 | 0.036714
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0003035 | 5.5713e-07 | 9 | 55004.4923 | 0.030645 |  | alpha (/C) |  |  |  | gamma (/V)
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030452 | 7.4346e-07 | 9 | -27502.2461 | 0.020447 |  | 1.3165e-05 | 9.6252e-06 | 217 |  | 0 | 0 | inf
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030446 | 6.3571e-07 | 9 | -31179.6263 | 0.019821
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.58e-07 | 9 | 25001.7185 | 0.018951
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 5.9878e-07 | 9 | 28344.7481 | 0.016972
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0003 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.1886e-08 | 8 | -110008.9845 | 0.0024076
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 3.7702e-07 | 9 | -5000.1675 | 0.0018852
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093983 | 4.6224e-07 | 9 | 3677.3801 | 0.0016998
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 89411.6197 | 0.0016988
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.9264e-07 | 9 | 2500.0838 | 0.0012316
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.7347e-07 | 9 | 2500.0838 | 0.0011837
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 3.1889e-07 | 9 | -3343.0296 | 0.0010661
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110008.9845 | 0.00089107
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 79941.588 | 0.00039971
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029373 | 0.00029373
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8941 | 0.0028839 | 3 | 0.029373 | 8.4708e-05
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.1886e-08 | 8 | 3677.3801 | 8.0482e-05
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3677.3801 | 2.9787e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.2956e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 12:20:17 | 20.7921 | 100003.6358 | 0.21873 | 87 | 0.43476 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6628 | 0.21001
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030413 | 7.3749e-07 | 9 | 55002.5095 | 0.040564
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 5.1851e-07 | 9 | -50003.5019 | 0.025928
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0003045 | 9.2626e-07 | 9 | -27501.2547 | 0.025473
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.4657e-07 | 9 | 25001.751 | 0.018666
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0003052 | 5.2179e-07 | 9 | -31184.5056 | 0.016272
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 3.5549e-07 | 9 | 28350.2425 | 0.010078
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.8706e-08 | 8 | -110005.019 | 0.0031578
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093989 | 7.5929e-07 | 9 | 3683.2508 | 0.0027966
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 5.625e-07 | 9 | -3348.4916 | 0.0018835
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 3.593e-07 | 9 | -5000.1707 | 0.0017965
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 89161.9208 | 0.0016941
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 5.1713e-07 | 9 | 2500.0853 | 0.0012929
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110005.019 | 0.00089104
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.0948e-07 | 9 | 2500.0853 | 0.00077373
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 79495.654 | 0.00039748
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029386 | 0.00029386
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.8706e-08 | 8 | 3683.2508 | 0.00010573
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8916 | 0.0024975 | 3 | 0.029386 | 7.3393e-05
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3683.2508 | 2.9834e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.5726e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 12:47:51 | 20.8246 | 100003.6621 | 0.21735 | 85 | 0.43212 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6891 | 0.21001
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030431 | 5.218e-07 | 9 | 55003.2593 | 0.028701
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 5.4757e-07 | 9 | -50003.5277 | 0.02738
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030492 | 7.5731e-07 | 9 | -31186.9666 | 0.023618
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.0018e-07 | 9 | 28352.1081 | 0.019851
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030512 | 6.3054e-07 | 9 | -27501.6296 | 0.017341
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 6.355e-07 | 9 | 25001.7639 | 0.015889
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 9.1065e-07 | 9 | -3350.3442 | 0.003051
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 5.9233e-07 | 9 | -5000.1719 | 0.0029618
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093989 | 6.921e-07 | 9 | 3685.337 | 0.0025506
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 89911.2208 | 0.0017083
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.2515e-07 | 9 | 2500.086 | 0.0010629
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.0343e-07 | 9 | 2500.086 | 0.0010086
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110006.5186 | 0.00089105
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 7.7815e-09 | 8 | -110006.5186 | 0.00085602
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 80837.3804 | 0.00040419
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029346 | 0.00029346
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3685.337 | 2.9851e-05
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 7.7815e-09 | 8 | 3685.337 | 2.8678e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8991 | 0 | 3 | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.8119e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 13:15:25 | 20.8196 | 100003.6629 | 0.21925 | 88 | 0.4357 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6899 | 0.21001
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030448 | 5.9097e-07 | 9 | 55007.0018 | 0.032508
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 5.4282e-07 | 9 | -50003.5297 | 0.027143
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030495 | 7.7584e-07 | 9 | -31190.2671 | 0.024199
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 1.8836e-07 | 8 | -110014.0035 | 0.020722
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.1355e-07 | 9 | 28353.1805 | 0.020231
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030506 | 7.0195e-07 | 9 | -27503.5009 | 0.019306
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.2687e-07 | 9 | 25001.7649 | 0.018173
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093983 | 5.535e-07 | 9 | 3686.7663 | 0.0020406
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 6.0474e-07 | 9 | -3351.4156 | 0.0020267
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 89161.9449 | 0.0016941
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 5.6858e-07 | 9 | 2500.086 | 0.0014215
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.8661e-07 | 9 | 2500.086 | 0.0012166
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 2.081e-07 | 9 | -5000.1721 | 0.0010405
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110014.0035 | 0.00089111
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 1.8836e-07 | 8 | 3686.7663 | 0.00069444
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 79495.6756 | 0.00039748
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029386 | 0.00029386
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8916 | 0.0024975 | 3 | 0.029386 | 7.3393e-05
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3686.7663 | 2.9863e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 2.4927e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 13:42:59 | 20.7946 | 100003.6182 | 0.21939 | 88 | 0.436 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6452 | 0.21001
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 7.8901e-07 | 9 | -50003.4859 | 0.039453
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030443 | 5.9264e-07 | 9 | 55003.0512 | 0.032597
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030478 | 7.6309e-07 | 9 | -31185.9085 | 0.023798
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030469 | 6.6559e-07 | 9 | -27501.5256 | 0.018305
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 6.1303e-07 | 9 | 28351.2296 | 0.01738
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 4.7052e-07 | 9 | 25001.743 | 0.011764
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 3.5858e-08 | 8 | -110006.1025 | 0.0039446
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093985 | 6.7224e-07 | 9 | 3684.3828 | 0.0024768
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 4.814e-07 | 9 | -5000.1699 | 0.0024071
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 88912.1462 | 0.0016893
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 4.8705e-07 | 9 | -3349.4866 | 0.0016314
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 5.8786e-07 | 9 | 2500.085 | 0.0014697
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 5.474e-07 | 9 | 2500.085 | 0.0013685
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110006.1025 | 0.00089105
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 79050.9001 | 0.00039525
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.0294 | 0.000294
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 3.5858e-08 | 8 | 3684.3828 | 0.00013211
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3684.3828 | 2.9844e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8891 | 0 | 3 | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.533e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 14:10:33 | 20.7946 | 100003.6673 | 0.22134 | 91 | 0.43966 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6943 | 0.21001
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0003045 | 5.9823e-07 | 9 | 55004.8899 | 0.032906
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 6.4054e-07 | 9 | -50003.5343 | 0.032029
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0003047 | 8.3014e-07 | 9 | -31187.5978 | 0.02589
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.1672e-07 | 8 | -110009.7799 | 0.023841
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 9.3526e-07 | 9 | 25001.7671 | 0.023383
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 8.1287e-07 | 9 | 28351.845 | 0.023046
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030485 | 7.3558e-07 | 9 | -27502.445 | 0.02023
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 8.2486e-07 | 9 | -3350.0779 | 0.0027634
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 4.6964e-07 | 9 | -5000.1723 | 0.0023483
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093986 | 6.0389e-07 | 9 | 3685.1528 | 0.0022254
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 88912.1899 | 0.0016893
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.6048e-07 | 9 | 2500.0862 | 0.0011512
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.6152e-07 | 9 | 2500.0862 | 0.00090384
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110009.7799 | 0.00089108
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 2.1672e-07 | 8 | 3685.1528 | 0.00079864
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 79050.9389 | 0.00039525
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.0294 | 0.000294
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3685.1528 | 2.985e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8891 | 0 | 3 | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.8943e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 14:38:07 | 20.7921 | 100003.6133 | 0.22034 | 89 | 0.43783 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6403 | 0.21001
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 8.3695e-07 | 9 | -50003.4814 | 0.04185
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030437 | 7.2153e-07 | 9 | 55007.0322 | 0.039689
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 8.4557e-07 | 9 | 28350.3224 | 0.023972
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030464 | 5.5503e-07 | 9 | -27503.5161 | 0.015265
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030459 | 3.8377e-07 | 9 | -31187.1705 | 0.011969
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 4.4525e-07 | 9 | 25001.7407 | 0.011132
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 4.4739e-07 | 9 | -5000.1697 | 0.002237
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 5.0489e-07 | 9 | -3348.5817 | 0.0016907
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 88662.3831 | 0.0016846
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093982 | 4.038e-07 | 9 | 3683.6544 | 0.0014875
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.2742e-07 | 9 | 2500.0849 | 0.0010686
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110014.0644 | 0.00089111
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.384e-07 | 9 | 2500.0849 | 0.00084602
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 4.0671e-09 | 8 | -110014.0644 | 0.00044744
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 78607.4039 | 0.00039304
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029413 | 0.00029413
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8866 | 0.0024975 | 3 | 0.029413 | 7.346e-05
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3683.6544 | 2.9838e-05
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 4.0671e-09 | 8 | 3683.6544 | 1.4982e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.2766e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 15:05:41 | 20.7671 | 100003.6108 | 0.21603 | 84 | 0.42962 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6378 | 0.21001
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030464 | 9.4841e-07 | 9 | -27502.2389 | 0.026084
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 4.6451e-07 | 9 | -50003.4796 | 0.023227
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.2059e-07 | 9 | 28348.0747 | 0.020427
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.000304 | 3.5423e-07 | 9 | 55004.4779 | 0.019484
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 7.0761e-07 | 9 | 25001.7398 | 0.017691
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030454 | 3.8355e-07 | 9 | -31183.2509 | 0.01196
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 5.3251e-08 | 8 | -110008.9557 | 0.0058581
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093984 | 8.4123e-07 | 9 | 3681.012 | 0.0030966
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 5.1217e-07 | 9 | -5000.1696 | 0.0025609
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 6.2834e-07 | 9 | -3346.3349 | 0.0021026
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 87913.1047 | 0.0016703
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 6.1567e-07 | 9 | 2500.0848 | 0.0015392
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 5.3141e-07 | 9 | 2500.0848 | 0.0013286
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110008.9557 | 0.00089107
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 77284.4104 | 0.00038642
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029454 | 0.00029454
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 5.3251e-08 | 8 | 3681.012 | 0.00019602
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3681.012 | 2.9816e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8791 | 0 | 3 | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 2.5713e-07 | 9 | 0 | 0
Ti612 100k | 10 | 03/03/2017 15:33:16 | 20.7646 | 100003.5902 | 0.22151 | 91 | 0.44003 | 3458A_518_VRC10ton1 | -2.7e-07 | 2.1e-06 | 75 | 100003.6172 | 0.21001
[8 empty] | Vd_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.00030378 | 7.6338e-07 | 9 | 55005.5703 | 0.04199
[8 empty] | V2_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 1.0003 | 6.951e-07 | 9 | -50003.459 | 0.034758
[8 empty] | V2_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 9.1961e-07 | 9 | 28345.6193 | 0.026067
[8 empty] | V2_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0003 | 8.9769e-07 | 9 | 25001.7295 | 0.022444
[8 empty] | Vd_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.00030421 | 7.7734e-07 | 9 | -27502.7852 | 0.021379
[8 empty] | Vd_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0003045 | 5.1132e-07 | 9 | -31181.182 | 0.015944
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 6.5264e-08 | 8 | -110011.1406 | 0.0071798
[8 empty] | Rlink HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0.0051171 | 0.00041536 | 97 | 10.0004 | 0.0041538
[8 empty] | SR104r_R0_LV <source?> | 9999.9993 | 0.00033 | 16 | 10.0004 | 0.0033001
[8 empty] | V2_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -1.0103 | 8.2194e-07 | 9 | -3343.8898 | 0.0027485
[8 empty] | V1_1 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -10 | 3.7879e-07 | 9 | -5000.1686 | 0.001894
[8 empty] | Vd_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | -0.0093981 | 4.7101e-07 | 9 | 3678.3968 | 0.0017326
[8 empty] | SR104r_alpha <source?> | 3.42e-07 | 1.9e-08 | 8 | 87913.0866 | 0.0016703
[8 empty] | V1_2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.249e-07 | 9 | 2500.0843 | 0.0010623
[8 empty] | V1_0 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 4.0994e-07 | 9 | 2500.0843 | 0.0010249
[8 empty] | 3458A_452_Vlin_Vd | 8e-09 | 8.1e-09 | 18 | -110011.1406 | 0.00089109
[8 empty] | SR104r_beta <source?> | -2.7e-08 | 5e-09 | 8 | 77284.3944 | 0.00038642
[8 empty] | T_def2 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0.01 | 3 | 0.029454 | 0.00029454
[8 empty] | Vdrift_gain HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 6.5264e-08 | 8 | 3678.3968 | 0.00024007
[8 empty] | 3458A_452_Vlin_gain | 8e-09 | 8.1e-09 | 18 | 3678.3968 | 2.9795e-05
[8 empty] | SR104r_Tref_LV <source?> | 20 | 0 | inf | 0 | 0
[8 empty] | SR104r_Vref_LV <source?> | 5 | 0 | inf | 0 | 0
[8 empty] | SR104r_gamma <source?> | 0 | 0 | inf | 0 | 0
[8 empty] | T2_av_gmh HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.8791 | 0 | 3 | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 3.833e-07 | 9 | 0 | 0
Processed with HRBA v1.3 on Thursday, 14. September 2017 02:20PM
|  | Run Id: | HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 |  |  |  |  | Uncertainty Budget | [6 empty] | R1(T) |  |  | exp. U(95%) | av T |  |  | av date/time | av V
Name | Test V | Date | T | R1 | std u. | dof | exp. U(95%) | Quantity (Label) | Value | Std u | dof | sens. coef. | u contrib. | LV | 1000070.713 | 2.9162 | 208 | 5.7492 | 20.7853 | 0.0060033 | 10 | 02/08/2017 18:01:46 | 1 | 4.2992e-08 | 72
Ti611 1M | 10.0001 | 02/08/2017 16:56:02 | 20.8045 | 1000068.825 | 2.2987 | 107 | 4.5571 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112 | HV | 1000068.6786 | 2.2599 | 100 | 4.4834 | 20.792 | 0.0060033 | 10 | 02/08/2017 17:39:51 | 10.0001 | 1.259e-07 | 43
| 1.9742e-07 |  | 0.010416 | [5 empty] | 1 | 7.4917e-07 | 150 | 1000069.075 | 0.74922
| 14.1495 |  | 3.5069 |  |  |  |  | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.99997 | 6.5225e-07 | 9 | -499609.2199 | 0.32587 |  | alpha (/C) |  |  |  | gamma (/V)
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.2176e-06 | 4.712e-07 | 9 | 549132.1532 | 0.25875 |  | -3.1129e-07 | 4.8609e-05 | 36 |  | -2.2603e-07 | 2.6507e-07 | 267
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 5.0101e-07 | 6.6913e-07 | 9 | -274566.0766 | 0.18372
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 5.8505e-07 | 9 | 249804.6099 | 0.14615
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 5.7909e-07 | 9 | 249804.6099 | 0.14466
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 2.4541e-07 | 2.6007e-07 | 9 | -274566.0766 | 0.071407
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.8715 | 0.0048082 | 3 | 8.9506 | 0.043037
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -1098264.3064 | 0.020867
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -10.0001 | 3.4851e-07 | 9 | -50003.0867 | 0.017427
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 81072.5988 | 0.013782
[8 empty] | Rlink HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.0063032 | 0.0013649 | 97 | 10.0003 | 0.01365
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.1468e-07 | 9 | 25001.5434 | 0.0078674
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 6.0145e-09 | 8 | -1098264.3064 | 0.0066056
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 1.9681e-07 | 9 | 25001.5434 | 0.0049206
[8 empty] | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.00090798 | 6.7326e-07 | 9 | -962.0928 | 0.00064774
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.001 | 6.2867e-07 | 9 | 875.3274 | 0.00055029
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -962.0928 | 1.828e-05
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 6.0145e-09 | 8 | -962.0928 | 5.7866e-06
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.4398e-07 | 9 | 0 | 0
Ti611 1M | 1 | 02/08/2017 17:17:57 | 20.7995 | 1000070.5884 | 3.0685 | 174 | 6.0561 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112
| 8.097e-08 |  | 0.010416 | [5 empty] | 1 | 1.5811e-06 | 178 | 1000072.1885 | 1.5813
| 26.3578 |  | 3.5069 |  |  |  |  | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.4299e-05 | 1.9337e-07 | 9 | -5614524.502 | 1.0857
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.728e-05 | 3.0718e-07 | 9 | 2646903.5366 | 0.81309
[8 empty] | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.67e-05 | 3.669e-07 | 9 | -1323451.7683 | 0.48557
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.667e-05 | 3.5245e-07 | 9 | -1323451.7683 | 0.46645
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.10005 | 5.0619e-08 | 9 | 5147490.68 | 0.26056
[8 empty] | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.099947 | 6.0051e-08 | 9 | -2426725.7683 | 0.14573
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 2.2385e-08 | 8 | -5614524.502 | 0.12568
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 2.2385e-08 | 8 | -5293807.0733 | 0.1185
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -5614524.502 | 0.10668
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -5293807.0733 | 0.10058
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 6.3113e-08 | 9 | 1213362.8841 | 0.076578
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 4.4951e-08 | 9 | 1213362.8841 | 0.054542
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.9802e-07 | 9 | 250014.7578 | 0.049508
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1 | 9.801e-08 | 9 | -500029.5155 | 0.049008
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.6511e-07 | 9 | 250014.7578 | 0.041279
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.8665 | 0.0025402 | 3 | 8.9506 | 0.022736
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 76072.3962 | 0.012932
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.4083e-07 | 9 | 0 | 0
Ti611 1M | 10.0001 | 02/08/2017 17:39:51 | 20.7945 | 1000068.7303 | 2.3157 | 109 | 4.5895 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112
| 2.0675e-07 |  | 0.010416 | [5 empty] | 1 | 7.4917e-07 | 150 | 1000068.9803 | 0.74922
| 18.3118 |  | 3.5069 |  |  |  |  | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.99997 | 7.6843e-07 | 9 | -499631.0814 | 0.38393
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.0068e-06 | 4.7686e-07 | 9 | 549016.1694 | 0.2618
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 8.2903e-07 | 9 | 249815.5407 | 0.2071
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 7.1816e-07 | 7.1875e-07 | 9 | -274508.0847 | 0.1973
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 1.2845e-07 | 8 | -1098032.3388 | 0.14105
[8 empty] | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 2.9232e-07 | 3.9801e-07 | 9 | -274508.0847 | 0.10926
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 2.3876e-07 | 9 | 249815.5407 | 0.059646
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.849 | 0.0041072 | 3 | 8.9506 | 0.036762
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -1098032.3388 | 0.020863
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -10.0001 | 3.3499e-07 | 9 | -50003.0826 | 0.01675
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 58571.0743 | 0.0099571
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.7085e-07 | 9 | 25001.5413 | 0.0092718
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.1235e-07 | 9 | 25001.5413 | 0.0078091
[8 empty] | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.00090842 | 7.1812e-07 | 9 | -913.83 | 0.00065624
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.001 | 4.2414e-07 | 9 | 831.6292 | 0.00035273
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 1.2845e-07 | 8 | -913.83 | 0.00011738
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -913.83 | 1.7363e-05
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 4.7065e-07 | 9 | 0 | 0
Ti611 1M | 1 | 02/08/2017 18:01:45 | 20.7845 | 1000070.4789 | 4.0398 | 43 | 8.1493 | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.4214e-05 | 4.7717e-07 | 9 | -5624138.6274 | 2.6837
| 6.584e-08 |  | 0.010416 |  |  |  |  | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112
| 20.3181 |  | 3.5069 | [5 empty] | 1 | 1.5811e-06 | 178 | 1000072.0791 | 1.5813
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.7169e-05 | 4.22e-07 | 9 | 2650001.3096 | 1.1183
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 8.6902e-08 | 8 | -5624138.6274 | 0.48875
[8 empty] | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.6513e-05 | 3.5516e-07 | 9 | -1325000.6548 | 0.47059
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 8.6902e-08 | 8 | -5300002.6191 | 0.46058
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.6821e-05 | 2.4588e-07 | 9 | -1325000.6548 | 0.32579
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.10005 | 4.7959e-08 | 9 | 5148842.5897 | 0.24693
[8 empty] | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.099947 | 6.2059e-08 | 9 | -2426049.6602 | 0.15056
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -5624138.6274 | 0.10686
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 8.6648e-08 | 9 | 1213024.8301 | 0.10511
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -5300002.6191 | 0.1007
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1 | 1.0173e-07 | 9 | -500029.443 | 0.050869
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 3.6085e-08 | 9 | 1213024.8301 | 0.043773
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.3265e-07 | 9 | 250014.7215 | 0.033164
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.018e-07 | 9 | 250014.7215 | 0.025452
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.8415 | 0.0025402 | 3 | 8.9506 | 0.022736
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 51070.6559 | 0.008682
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.9273e-07 | 9 | 0 | 0
Ti611 1M | 10.0001 | 02/08/2017 18:23:40 | 20.777 | 1000068.4809 | 2.2814 | 104 | 4.5242 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112
| 2.4687e-07 |  | 0.010316 | [5 empty] | 1 | 7.4917e-07 | 150 | 1000068.7309 | 0.74922
| 13.8644 |  | 3.3837 |  |  |  |  | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 2.626e-07 | 7.7866e-07 | 9 | -274361.9361 | 0.21363
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.2156e-06 | 3.711e-07 | 9 | 548723.8723 | 0.20363
[8 empty] | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.99997 | 3.3087e-07 | 9 | -499607.0928 | 0.16531
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 5.0445e-07 | 5.8689e-07 | 9 | -274361.9361 | 0.16102
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 6.0958e-07 | 9 | 249803.5464 | 0.15227
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.99997 | 4.5089e-07 | 9 | 249803.5464 | 0.11263
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.834 | 0.0029216 | 3 | 8.9506 | 0.02615
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -10.0001 | 4.387e-07 | 9 | -50003.0725 | 0.021936
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -1097447.7445 | 0.020852
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.8052e-07 | 9 | 25001.5363 | 0.0095136
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 43570.0471 | 0.0074069
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 2.4598e-07 | 9 | 25001.5363 | 0.0061499
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 3.2285e-09 | 8 | -1097447.7445 | 0.0035431
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1.001 | 7.5046e-07 | 9 | 879.2887 | 0.00065987
[8 empty] | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.00090863 | 4.6182e-07 | 9 | -965.7323 | 0.00044599
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -965.7323 | 1.8349e-05
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 3.2285e-09 | 8 | -965.7323 | 3.1178e-06
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 10.0001 | 3.536e-07 | 9 | 0 | 0
Ti611 1M | 1 | 02/08/2017 18:45:35 | 20.772 | 1000070.9833 | 3.3611 | 75 | 6.6959 | Ti612_R0_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 100003.6255 | 0.21111 | 77 | 10.0003 | 2.1112
| 7.5787e-08 |  | 0.010316 |  |  |  |  | Vd_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.4298e-05 | 3.3519e-07 | 9 | -5617138.5378 | 1.8828
| 25.3993 |  | 3.3837 | [5 empty] | 1 | 1.5811e-06 | 178 | 1000072.5834 | 1.5813
[8 empty] | Vd_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -4.7058e-05 | 1.9211e-07 | 9 | 2653948.374 | 0.50985
[8 empty] | Vd_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.667e-05 | 3.0109e-07 | 9 | -1326974.187 | 0.39954
[8 empty] | Vd_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 4.6731e-05 | 2.8193e-07 | 9 | -1326974.187 | 0.37411
[8 empty] | V2_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.10005 | 7.1994e-08 | 9 | 5142019.416 | 0.37019
[8 empty] | V2_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0.099947 | 8.5214e-08 | 9 | -2429467.2414 | 0.20703
[8 empty] | 3458A_452_Vlin_gain | 0 | 1.9e-08 | 18 | -5617138.5378 | 0.10673
[8 empty] | 3458A_452_Vlin_Vd | 0 | 1.9e-08 | 18 | -5307896.748 | 0.10085
[8 empty] | T_def2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 0.01 | 3 | 8.9506 | 0.089506
[8 empty] | Vdrift_gain HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 1.3378e-08 | 8 | -5617138.5378 | 0.075145
[8 empty] | Vdrift_Vd HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 0 | 1.3378e-08 | 8 | -5307896.748 | 0.071008
[8 empty] | V2_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 5.4902e-08 | 9 | 1214733.6207 | 0.066691
[8 empty] | V2_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -0.099946 | 5.2073e-08 | 9 | 1214733.6207 | 0.063255
[8 empty] | V1_1 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | -1 | 1.0185e-07 | 9 | -500029.7231 | 0.05093
[8 empty] | V1_0 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.6012e-07 | 9 | 250014.8616 | 0.040033
[8 empty] | V1_2 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.5736e-07 | 9 | 250014.8616 | 0.039342
[8 empty] | T2_av_gmh HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 20.8265 | 0.0025402 | 3 | 8.9506 | 0.022736
[8 empty] | Ti612_TRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 20.7904 | 0.0014778 | 13 | -8.9506 | 0.013228
[8 empty] | Ti612_alpha_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 8.95e-06 | 1.7e-07 | 8 | 36069.6296 | 0.0061318
[8 empty] | Ti612_VRef_LV_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 10 | 9.1808e-08 | 147 | 0 | 0
[8 empty] | Ti612_beta_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | Ti612_gamma_HRBC.v0.4 Ti612 100k:SR104r 10k 03/03/2017 11:37:33 | 0 | 0 | inf | 0 | 0
[8 empty] | V1_3 HRBC.v0.4 Ti611 1M:Ti612 100k 02/08/2017 16:01:48 | 1 | 1.6752e-07 | 9 | 0 | 0